National Metrological Verification Regulation of People's Republic of China

JJG 11482022

AC Charging Piles for Electric Vehicles (for Trial Implementation)

Issued on 20220628                   Enforced from 20221228

Issued by State Administration for Market Regulation

Verification Regulation of AC Charging Piles for Electric Vehicles

(for Trial Implementation)

JJG 1148-2022

Superseding

JJG 1148-2018

Competent Authority:

National Electromagnetic Metrology Technical Committee

Main Drafting Organizations:

National Institute of Metrology (NIM)

Shenzhen Academy of Metrology & Quality Inspection (SMQ)

Participating Drafting Organizations:

Hangzhou Institute of Quality and Metrology (HIQM)

This regulation is entrusted to National Electromagnetic Metrology Technical Committee for interpretation.

Primary Drafters:

HUANG Hongtao (National Institute of Metrology)

CHEN Shaohui (Shenzhen Academy of Metrology & Quality Inspection)

WANG Lei (National Institute of Metrology)

Participating Drafters:

CHEN Ruiheng (Shenzhen Academy of Metrology & Quality Inspection)

LIU Lijuan (National Institute of Metrology)

FAN Junjie (Hangzhou Institute of Quality and Metrology)

Table of Contents

Introduction	I

1	Scope	1

2	Normative References	1

3	Terms and Definitions	1

4	Overview	2

5	Metrological Performance Requirements	3

5.1 Operational Error	3

5.2 Clock Time Error	3

6	General Technical Requirements	4

6.1 Markings	4

6.2 Test Interface	4

6.3 Constant	4

6.4 Minimum Energy Increment	4

6.5 Display of Delivered Energy	4

7	Control of Measuring Instruments	4

7.1 Initial Verification, Subsequent Verification, and In-Service Inspection	4

7.2 Verification Conditions	5

7.3 Verification Items and Verification Methods	6

7.4 Processing of Verification Results	10

7.5 Verification Period	10

Appendix A - Verification Original Record Format	11

Appendix B - Format of Inner Page of Verification Certificate / Verification Result Notification (Page 2)	14

Appendix C - Format of the Verification Result Page of Verification Certificate / Verification Result Notification (Page 3)	15

Introduction

JJF 1002-2010 The Rules for Drafting National Metrological Verification Regulation, JJF 1001-2011 General Terms in Metrology and Their Definitions and JJF 1059.1-2012 Evaluation and Expression of Uncertainty in Measurement together form the fundamental series of standards that support the development of this regulation.

This regulation is a revision of JJG 1148-2018 AC Charging Spot for Electric Vehicles. Compared with JJG 1148-2018, apart from editorial changes, the significant technical changes are as follows:

——Revised the scope of application of the regulation.

——Added the definition of connection points and revised the definition of AC charging piles in the terminology clause.

——Removed the insulation resistance test, the transaction amount error and the indication error from the verification items.

——Added the insulation resistance requirements to the verification conditions.

——Adjusted the selection of test load points.

——Revised the verification method for operational errors.

——Adjusted the verification period of AC charging piles.

The previous editions of this regulation are as follows:

——JJG 1148-2018.

Verification Regulation of AC Charging Piles for Electric Vehicles

(for Trial Implementation)

Scope

This regulation applies to the initial verification, subsequent verification, and in-service inspection of electric vehicle AC charging piles.

The verification of AC charging piles based on statistical sampling plans may be conducted with reference to this regulation.

Normative References

This regulation refers to the following documents:

JJG 596 Verification Regulation of Fixed Alternating-current Electrical Energy Meters

JJG 597 Verification Regulation of Verification Equipment for AC Electrical Energy Meters

JJF 1139 Principle and Method for Determination of Verification Period of Measuring Instruments

For dated referenced documents, only the specified version shall apply to this regulation. For undated referenced documents, the latest version including all amendment sheets shall apply.

Terms and Definitions

The following terms and definitions apply to this regulation.

## 3.1 AC Charging Pile

A dedicated device that supplies AC electrical energy to electric vehicles equipped with on-board chargers via conductive means, and incorporates functions such as AC energy measurement, data storage, and provision of results to users.

## 3.2 Connection Point

The point at which an electric vehicle is connected to an AC charging pile.

Note: If the output cable is a fixed part of the AC charging pile, this point refers to the connector at the cable end. Otherwise, it refers to the point where the cable is inserted into the AC charging pile.

## 3.3 Real Load Test Method

The method for conducting verification with real operating loads.

## 3.4 Phantom Load Test Method

The method for verification using the phantom load provided by an AC charge pile verification device.

## 3.5 Operational error

AC energy measurement error of charging piles under operational conditions.

## 3.6 Test Output

The output of pulses corresponding to the AC energy measured by the charging pile.

## 3.7 Constant

A numerical value indicating the relationship between the electrical energy recorded by the charging pile and the corresponding test output.

## 3.8 Nominal Voltage

The output voltage specified by the manufacturer of the charging pile under normal operation.

Overview

An electric vehicle AC charging pile (hereinafter referred to as the charging pile) is a metering device that supplies AC electric energy to electric vehicles. It is mainly composed of a housing, a conductive charging connection assembly, a control unit, a metering module, and a display or data transmission interface. Its schematic diagram is shown in Figure 1.

The charging pile, controlled by the control unit, supplies the required AC electrical energy to the electric vehicle. The electric energy is delivered to the electric vehicle through the conductive charging connection assembly to meet the charging demand. The metering module performs the energy measurement and outputs the measured values to the display or data transmission interface.

Figure 1 Schematic Diagram of the AC Charging Piles

Metrological Performance Requirements

## 5.1 Operational Error

The operational error of a charging pile shall be expressed as a relative error. Under the reference verification conditions specified in 7.2.1, the operational error limits shall comply with Table 1. The average temperature coefficient for energy metering of charging piles is specified in Table 2.

Table 1 Operational Error Limits for Charging Piles

| Load Current I① | Power Factor cosφ② | Charging Pile Accuracy Class |  |
| --- | --- | --- | --- |
|  |  | Class 1 | Class 2 |
|  |  | Operational Error Limit③ |  |
| Imin ≤I ≤ Imax | 1 | ±1.0% | ±2.0% |
| Note： ① Imin—minimum current, Imax—maximum current. ② φ is the phase difference between the phase voltage and the phase current of the load branch. ③Under special ambient temperature conditions (-20℃ ≤ t < -10℃ or +40℃ < t ≤ +50℃), considering the influence of ambient temperature variation, the absolute value of the operational error limit shall be adjusted by adding a positive correction value e: e = C×¦⊿t¦ where, C—average temperature coefficient for energy metering of charging piles, %/℃; ⊿t—ambient temperature deviation value. For high temperatures, it is the difference between the current ambient temperature and +40°C. For low temperatures, it is the difference between the current ambient temperature and -10°C. The ambient temperature shall be measured at a distance of more than 0.5 m from the charging pile. |  |  |  |

Table 2 Average Temperature Coefficient for Energy Metering of Charging Piles

| Load Current I | Power Factor cosφ | Charging Pile Accuracy Class |  |
| --- | --- | --- | --- |
|  |  | Class 1 | Class 2 |
|  |  | Average Temperature Coefficient (%/℃) |  |
| Imin ≤ I ≤ Imax | 1 | 0.05 | 0.10 |

## 5.2 Clock Time Error

For charging piles with time-of-use billing functionality, the clock time error shall not exceed  5 s for initial verification, and 3 min for subsequent verification.

General Technical Requirements

## 6.1 Markings

The nameplate shall contain the following markings:

—Name and model;

—Manufacturer's name;

—Standards with which the product complies;

—Serial number and year of manufacture;

—Nominal voltage, minimum current, and maximum current①;

—Constant;

—Accuracy class②;

—Unit of measurement (which can be displayed on the display screen).

Note:

① For charging piles installed before January 1, 2023, if the maximum and minimum currents are not marked on the charging pile, the maximum current shall be calculated as the rated power divided by the nominal voltage; the minimum current shall be taken as Imax/10 or not less than 6 A.

② For charge piles installed before 1 January 2023 without an accuracy class marking, class 2 shall apply.

## 6.2 Test Interface

The technical specifications of the test output and communication interfaces provided by the charging pile for verification purposes shall comply with the relevant standards.

## 6.3 Constant

Charging piles shall have pulse output corresponding to the AC energy measured by the charging pile, and the pulse constant shall consistent with to that marked on the nameplate.

## 6.4 Minimum Energy Increment

The minimum energy increment shall be 0.001 kWh.

## 6.5 Display of Delivered Energy

Charging piles shall be capable of displaying the delivered energy. The energy display shall have no fewer than 6 digits (including at least 3 decimal places). For charging piles equipped with time-of-use billing functionality, the display resolution of the clock time shall not exceed 1 s.

Control of Measuring Instruments

## 7.1 Initial Verification, Subsequent Verification, and In-Service Inspection

Initial verification refers to the verification of charging piles that have not been verified before. Subsequent verification refers to any verification carried out after the initial verification. Charging piles that have been repaired shall undergo initial verification. In-service inspection is a type of examination to determine whether the verification mark or verification certificate of the charging pile remains valid and whether the charging pile has been significantly altered after verification.

## 7.2 Verification Conditions

7.2.1 Verification Environmental Conditions

When verifying charging piles, the following conditions shall be met:

The verification conditions and their permissible deviations shall not exceed the limits specified in Table 3.

During the verification of charging piles, there shall be no significant impact loads, and the seals of charging piles shall be intact.

Table 3 Verification Conditions and Permissible Deviations

| Item | Requirement | Remark |
| --- | --- | --- |
| Ambient Temperature | -20°C ~ +50°C | / |
| Relative Humidity | ≤ 90% | / |
| Permissible deviation of Output Current① | ±5% | / |
| Power Factor of Charging Pile Output | ≥ 0.90 | / |
| Voltage and Current Waveform of Charging Pile Output | Total Harmonic Distortion ≤ 5% | / |
| Insulation Resistance | ≥ 5 MΩ | Between live circuits without electrical connection, and between each independent live circuit and ground (metal enclosure), when the normal operating voltage of the charging sopt ≤ 300 V, the insulation test voltage is 500 V; when the normal operating voltage of the charging pile > 300 V, the insulation test voltage is 1 kV. |
| ① Deviation between the actual output current and the specified verification current value. |  |  |

7.2.2 Measurement Standard Instrument and Main Auxiliary Equipment

7.2.2.1 Verification Device and Load

When using the phantom load verification method or the real load verification method, the verification device shall comply with the relevant requirements of JJG 597. The load may be either an electronic load or a physical load, and its impedance and power shall be adjustable.

7.2.2.2 Thermometer

The measurement range shall cover -25 °C to +55 °C, and the maximum permissible error of temperature measurement shall not exceed ±1 ℃.

7.2.2.3 Reference Clock

The absolute value of the time error of the reference clock shall not exceed 1 s.

## 7.3 Verification Items and Verification Methods

7.3.1 Verification Items

The verification items for charging piles are listed in Table 4.

Table 4 Charging Pile Verification Items

| Verification Items | Initial Verification | Subsequent Verification | In-Service Inspection |
| --- | --- | --- | --- |
| Inspection of Appearance and Function | + | + | + |
| Operational Error | + | + | − |
| Clock Time Error | + | + | − |
| Note: “+” denotes mandatory verification items, “−” denotes optional verification items. |  |  |  |

7.3.2 Verification Methods

7.3.2.1 Inspection of Appearance and Function

In addition to complying with the requirements of clauses 6.1 and 6.2 of this regulation, the appearance of charging piles shall also be inspected for the following items:

a) Whether the markings are complete and legible;

b) Whether there is any obvious damage;

c) Whether there are measures to prevent unauthorized persons from inputting data or operating;
  d) Whether the number of display digits complies with the requirements;
  e) Whether the basic functions are normal.

Note: For charging piles installed before January 1, 2023, compliance with this item may be deferred until December 31, 2027.

7.3.2.2 Operational Error

a) Selection of Test Load Points

When determining the operational error of a charging pile, a load point shall be selected within the range from 50%Imax to Imax . Additional load points are permitted as needed.

b) Verification of Operational Error

The operational error of the charging pile under test is determined by comparing the energy values measured simultaneously by the verification device and the charging pile under test.

Based on principles, the verification method can be categorized as the real load verification method and the phantom load verification method.

1) Real Load Verification Method

The verification device, composed of a charging pile calibrator and a load, is connected to the charging connection point. Under the selected load point, the operational error of the charging pile under test is determined by comparing the energy values measured simultaneously by the verification device and the charging pile under test using method I or method II.

Method I:

With both the charging pile calibrator and the charging pile under test operating continuously, the operating error of the charging pile under test is determined by using the pulses output from the charging pile under test to control the counting process of the calibrator. When this method is used, method II shall be applied to check whether the constant of the charging pile is correct.

The operational error  of the charging pile under test is calculated according to formula (1),

(1)

where,

m — measured pulse count;

m0 — calculated pulse count, is calculated according to formula (2),

(2)

where,

N — Pulse count of the charging pile under test;

C0 — Constant of the charging pile calibrator, imp/kWh;

CL — Constant of the charging pile under test, imp/kWh.

Note: The pulse count N of the charging pile under test shall be appropriately selected so that the quantization error of the calculated pulse count is not greater than 1/10 of the class index of the charging pile under test.

Method II:

With both the charging pile calibrator and the charging pile under test operating continuously, the operational error of the charging pile under test is determined by comparing the energy indication value measured by the charging pile under test with the energy value measured by the charging pile calibrator.

The operational error  of the charging pile under test is calculated according to formula (3),

(3)

where,

— The difference between the energy indication value of the charging pile under test at the stop of charging and that at the start of charging, kWh;

E — Energy value measured by the charging pile calibrator, kWh;

— The determined systematic error of the verification device. When no correction is required, ＝0.

The calibrator shall operate synchronously with the charging pile under test. The ratio (%) of the energy value represented by the least significant digit (or the minimum scale division) of the display of the charging pile under test to the accumulated energy value  shall not exceed 1/5 of the class index of the charging pile under test.

During the verification process, error measurement data shall be recorded at least twice, and the operational error shall be calculated as the average of the recorded values.

2) Phantom Load Verification Method

Using the verification device, the operational error of the charging pile under test is determined by comparing the energy values measured simultaneously by the verification device and the charging pile under test. When using the phantom load verification method, the influence of the connection position on the verification results shall be taken into account.

Verification of charging piles by the standard meter method shall be performed as specified in the provisions of 1).

Verification of charging piles by the watt-second method shall be performed as follows. The constant power value is determined either by measuring the set power value of a power source with a power meter, or by using the set power value of a standard power source. The quantization error due to the number of display digits of the power meter or the power source shall not exceed 1/10 of the class index of the charging pile under test. At the same time, a standard timer is used to measure the time for the charging pile to output a certain number of pulses under the constant power. The actual energy is obtained by multiplying the time by the constant power. The operational error of the charging pile under test is determined by comparing the energy measured by the charging pile and the actual energy.

The operational error  of the charging pile under test is calculated according to formula (4),

(4)

where,

m — measured pulse count;

m0 — calculated pulse count, calculated according to formula (5),

(5)

where,

CL — Constant of the charging pile under test, imp/kWh;

Tn — Measurement time, s;

P — Determined constant power value, W.

The charging pile under test shall operate continuously. The standard timer shall be controlled by an automatic method, and the measurement time shall be no less than 10 s. If the power is not sufficiently stable, the measurement time shall be increased appropriately.

During the verification process, error measurement data shall be recorded at least twice, and the operational error shall be calculated as the average of the recorded values.

7.3.2.3 Clock Time Error

Comparing the time displayed on the charging pile with that of the reference clock, the clock time error ΔT of the charging pile is calculated according to formula (6),

(6)

where,

T — Displayed time of the reference clock, ××h××min××s;

T0 —Displayed time of the charging pile under test, ××h××min××s.

The test results shall meet the requirements of clause 5.2.

- 7.4 Processing of Verification Results

7.4.1 Processing of Verification Results

7.4.1.1 All data shall be judged based on the rounded data.

7.4.1.2 The rounding interval for the operational error shall be 1/10 of the accuracy class of the charging pile.

7.4.1.3 The rounding interval for the clock time error shall be 1 s.

7.4.2 Rounding Method for Measurement Data

7.4.2.1 Rounding method when the rounding interval is 1: For the digit immediately to the right of the retained digit, if it is greater than 0.5 relative to the retained digit value of 1, add 1 to the retained digit; if it is less than 0.5, the retained digit remains unchanged; if it is equal to 0.5, the retained digit remains unchanged if the retained digit is even, and add 1 if it is odd.

Note: The "retained digit" refers to the digit corresponding to the rounding interval position, and this value is called the retained digit.

7.4.2.2 Rounding method when the rounding interval is n (n ≠ 1): Divide the measured data by n, then round the result according to the rounding method specified in clause 7.4.2.1. After rounding, multiply by n to obtain the final rounded result.

7.4.3 If all items meet the requirements, the charging pile is judged as qualified; otherwise, it is judged as unqualified. A qualified charging pile shall be issued a verification certificate (see the appendix B and C for the certificate format) and a verification pass mark.

7.4.4 A qualified charging pile shall be sealed at the position of the internal electrical energy meter or metering module.

7.4.5 An unqualified charging pile shall be issued a verification result notification, and the original verification pass seal or verification pass mark shall be invalidated.

## 7.5 Verification Period

The verification period of charging piles shall generally not exceed 3 years, and may be adjusted according to JJF 1139 and the requirements of the competent administrative department of metrology. For charging piles that are equipped with online monitoring of metrological performance or anti-cheating functions, and have implemented data transmission, the verification period may be appropriately extended with the approval of the competent administrative department of metrology.

Appendix A - Verification Original Record Format

AC Charging Pile Verification Original Record

Certificate Number:

| Submitted Instrument Information: |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Application Number: |  |  |  |  |  | Customer: |  |  |  |  |
| Instument Name: |  |  |  |  |  | Manufacturer: |  |  |  |  |
| Model/Type: |  |  |  |  |  | Serial Number: |  |  |  |  |
| Verification Basis: |  |  |  |  |  |  |  |  |  |  |
| Verification Environmental Conditions and Location: |  |  |  |  |  |  |  |  |  |  |
| Temperature | ℃ |  |  | Location: |  |  |  |  |  |  |
| Relative Humidity: | % |  |  | Other: |  |  |  |  |  |  |
| (Primary) Standard Measuring Instruments Used for Verification: |  |  |  |  |  |  |  |  |  |  |
| Name |  |  | Measuring Range |  | Uncertainty/Accuracy Class/Maximum Permissible Error |  |  |  | Certificate Number of (Primary) Standard Measuring Instruments | Valid Until |

Page   of   Pages

AC Charging Pile Verification Original Record (Method I)

Certificate Number:

| Verification Result Record |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Verification Item |  |  |  |  | Verification Result |  |  |  |  |  |  |  |  |
| 1. Appearance and Function Inspection |  |  |  |  | □Qualified / □Unqualified |  |  |  |  |  |  |  |  |
| 2. Operational Error |  |  |  |  | □Qualified / □Unqualified |  |  |  |  |  |  |  |  |
| Voltage | Current |  | Power Factor |  |  | Operational Error (%) Correction Value due to Temperature Variation: e = |  |  |  |  |  |  |  |
|  |  |  |  |  |  | Error 1 |  | Error 2 |  | Average Value |  | Rounded Value |  |
| Constant Check |  |  |  |  | □Qualified / □Unqualified |  |  |  |  |  |  |  |  |
| Operation under Applied Load |  | Indicated Energy Value of Charging Pile (kWh) |  |  |  |  |  |  |  | Reference Energy Value E (kWh) |  |  | Error (%) |
|  |  | E'before |  |  |  | E'after |  |  | E' |  |  |  |  |
| 3.Clock Time Error |  |  |  | □Qualified / □Unqualified |  |  |  |  |  |  |  |  |  |
| Clock Time Error |  |  |  | Displayed Time of Charging Pile |  |  | Displayed Time of Reference Clock |  |  |  | Error (s) |  |  |
| Verification Conclusion and Explanation |  |  |  | □Qualified / □Unqualified |  |  |  |  |  |  |  |  |  |

Verifier:         Reviewer:        Verification Date:    Year   Month   Day

Page   of   Pages

AC Charging Pile Verification Original Record (Method II)

Certificate Number:

| Verification Result Record |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Verification Item |  |  | Verification Result |  |  |  |  |  |  |  |  |
| 1. Appearance and Function Inspection |  |  | □Qualified / □Unqualified |  |  |  |  |  |  |  |  |
| 2. Operational Error |  |  | □Qualified / □Unqualified |  |  |  |  |  |  |  |  |
| Correction Value due to Temperature Variation: e = |  |  |  |  |  |  |  |  |  |  |  |
| Voltage: Current: Power Factor: |  |  |  |  |  |  |  |  |  |  |  |
| Operation under Applied Load | No. | Indicated Energy Value of Charging Pile (kWh) |  |  |  | Reference Energy Value (kWh) |  | Error (%) |  | Average Value (%) | Rounded Value (%) |
|  |  | E'before |  | E'after | E' | E |  |  |  |  |  |
|  | 1 |  |  |  |  |  |  |  |  |  |  |
|  | 2 |  |  |  |  |  |  |  |  |  |  |
| Voltage: Current: Power Factor: |  |  |  |  |  |  |  |  |  |  |  |
| Operation under Applied Load | No. | Indicated Energy Value of Charging Pile (kWh) |  |  |  | Reference Energy Value (kWh) |  | Error (%) |  | Average Value (%) | Rounded Value (%) |
|  |  | E'before |  | E'after | E' | E |  |  |  |  |  |
|  | 1 |  |  |  |  |  |  |  |  |  |  |
|  | 2 |  |  |  |  |  |  |  |  |  |  |
| 3.Clock Time Error |  |  | □Qualified / □Unqualified |  |  |  |  |  |  |  |  |
| Clock Time Error |  |  | Displayed Time of Charging Pile |  |  |  | Displayed Time of Reference Clock |  | Error (s) |  |  |
| Verification Conclusion and Explanation |  |  | □Qualified / □Unqualified |  |  |  |  |  |  |  |  |

Verifier:         Reviewer:        Verification Date:    Year   Month   Day

Page   of   Pages

Appendix B - Format of Inner Page of Verification Certificate / Verification Result Notification (Page 2)

Certificate Number:

| Authorization Statement for Verification Institution |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Verification Environmental Conditions and Location: |  |  |  |  |  |
| Temperature | ℃ | Location: |  |  |  |
| Relative Humidity | % | Other: |  |  |  |
| (Primary) Standard Measuring Instruments Used for Verification: |  |  |  |  |  |
| Name | Measuring Range | Uncertainty/Accuracy Class/Maximum Permissible Error |  | Certificate Number of (Primary) Standard Measuring Instruments | Valid Until |

Page   of   Pages

Appendix C - Format of the Verification Result Page of Verification Certificate / Verification Result Notification (Page 3)

C.1 Page 3 of Verification Certificate

Certificate Number XXXXXX-XXXX

Verification Result

1. Inspection of Appearance and Function: 2. Operational Error: Operational Error (%) under □Single-Phase □Three-Phase Balanced Load. Voltage; Current; cosφ = 1. 3. Clock time error: Verification Conclusion:

Page   of   Pages

C.2 Page 3 of Verification Result Notification

Certificate Number XXXXXX-XXXX

Verification Result

1. Inspection of Appearance and Function: 2. Operational Error: Operational Error (%) under □Single-Phase □Three-Phase Balanced Load. Voltage; Current; cosφ = 1. 3. Clock time error: Non-conforming Items of Verification Result:

Page   of   Pages